## Summary
HE4 Stab 11 — TP1 Summary Report
Study: |  | HE4 Stab 11
Timepoint: |  | TP1
Instruments: |  | AIM 4, AIM 5
Lots: |  | Lot1, Lot2, Validity
Run Date: |  | 21 May 2026
Lot1 — Run Summary
Instrument | Run ID | Sample | Reps | RLU Avg | Dose Avg (pmol/L) | RLU %CV | Dose %CV | Mean Spec Low | Mean Spec High | Result
AIM 4 | Run1 | Panel1 | 12 / 12 | 349386 | 45 | 3.024% | 2.75% | 41.4 | 48.2 | PASS
AIM 4 | Run1 | Panel2 | 12 / 12 | 1183139 | 154.1 | 3.771% | 3.149% | 145.6 | 169.2 | FAIL
AIM 4 | Run1 | Panel3 | 12 / 12 | 3433440 | 456.4 | 2.72% | 2.618% | 424.8 | 493.6 | PASS
AIM 4 | Run1 | Panel4 | 12 / 12 | 7369988 | 1121 | 2.476% | 2.196% | 1047.7 | 1217.5 | FAIL
AIM 4 | Run1 | RefCon1 | 12 / 12 | 288587 | 37.1 | 3.143% | 2.579% | 34.1 | 39.7 | FAIL
AIM 4 | Run1 | RefCon2 | 12 / 12 | 1192100 | 150.6 | 3.126% | 2.736% | 138.3 | 160.7 | FAIL
AIM 4 | Run1 | RefCon3 | 12 / 12 | 4738966 | 648.3 | 3.117% | 3.074% | 601.7 | 699.3 | FAIL
AIM 4 | Run2 | Panel1 | 12 / 12 | 349485 | 44.9 | 4.3% | 3.836% | 41.4 | 48.2 | FAIL
AIM 4 | Run2 | Panel2 | 12 / 12 | 1213263 | 156.2 | 2.722% | 2.477% | 145.6 | 169.2 | PASS
AIM 4 | Run2 | Panel3 | 12 / 12 | 3446366 | 459.8 | 2.07% | 2.237% | 424.8 | 493.6 | PASS
AIM 4 | Run2 | Panel4 | 12 / 12 | 7523870 | 1145.5 | 3.472% | 3.042% | 1047.7 | 1217.5 | FAIL
AIM 4 | Run2 | RefCon1 | 12 / 12 | 288467 | 36.8 | 4.056% | 3.556% | 34.1 | 39.7 | FAIL
AIM 4 | Run2 | RefCon2 | 12 / 12 | 1180225 | 150 | 3.312% | 3.163% | 138.3 | 160.7 | FAIL
AIM 4 | Run2 | RefCon3 | 12 / 12 | 4708026 | 649.5 | 2.397% | 2.332% | 601.7 | 699.3 | PASS
AIM 4 | Run3 | Panel1 | 12 / 12 | 343130 | 43.8 | 2.392% | 2.397% | 41.4 | 48.2 | PASS
AIM 4 | Run3 | Panel2 | 12 / 12 | 1205130 | 156.9 | 2.272% | 2.191% | 145.6 | 169.2 | PASS
AIM 4 | Run3 | Panel3 | 12 / 12 | 3456700 | 462 | 2.29% | 1.674% | 424.8 | 493.6 | PASS
AIM 4 | Run3 | Panel4 | 12 / 12 | 7478332 | 1134.9 | 3.193% | 3.22% | 1047.7 | 1217.5 | FAIL
AIM 4 | Run3 | RefCon1 | 12 / 12 | 291894 | 37.3 | 3.39% | 3.176% | 34.1 | 39.7 | FAIL
AIM 4 | Run3 | RefCon2 | 12 / 12 | 1184751 | 149.1 | 3.23% | 2.95% | 138.3 | 160.7 | FAIL
AIM 4 | Run3 | RefCon3 | 12 / 12 | 4763350 | 653.4 | 1.933% | 2.375% | 601.7 | 699.3 | PASS
AIM 5 | Run1 | Panel1 | 12 / 12 | 343362 | 43.9 | 2.512% | 1.898% | 41.4 | 48.2 | PASS
AIM 5 | Run1 | Panel2 | 12 / 12 | 1197396 | 155 | 3.003% | 2.999% | 145.6 | 169.2 | PASS
AIM 5 | Run1 | Panel3 | 12 / 12 | 3411116 | 454.2 | 2.915% | 2.27% | 424.8 | 493.6 | FAIL
AIM 5 | Run1 | Panel4 | 12 / 12 | 7415080 | 1126.6 | 3.22% | 2.493% | 1047.7 | 1217.5 | FAIL
AIM 5 | Run1 | RefCon1 | 12 / 12 | 290542 | 37.3 | 2.296% | 2.365% | 34.1 | 39.7 | PASS
AIM 5 | Run1 | RefCon2 | 12 / 12 | 1177697 | 148.5 | 4.098% | 3.83% | 138.3 | 160.7 | FAIL
AIM 5 | Run1 | RefCon3 | 12 / 12 | 4785178 | 653.5 | 2.659% | 2.422% | 601.7 | 699.3 | FAIL
AIM 5 | Run2 | Panel1 | 12 / 12 | 347408 | 44.4 | 4.184% | 3.866% | 41.4 | 48.2 | FAIL
AIM 5 | Run2 | Panel2 | 12 / 12 | 1219369 | 157.3 | 2.348% | 2.351% | 145.6 | 169.2 | PASS
AIM 5 | Run2 | Panel3 | 12 / 12 | 3413381 | 456.5 | 3.381% | 3.232% | 424.8 | 493.6 | FAIL
AIM 5 | Run2 | Panel4 | 12 / 12 | 7464346 | 1131.1 | 2.639% | 2.18% | 1047.7 | 1217.5 | FAIL
AIM 5 | Run2 | RefCon1 | 12 / 12 | 286083 | 36.6 | 3.002% | 2.285% | 34.1 | 39.7 | FAIL
AIM 5 | Run2 | RefCon2 | 12 / 12 | 1185646 | 150.1 | 3.472% | 3.458% | 138.3 | 160.7 | FAIL
AIM 5 | Run2 | RefCon3 | 12 / 12 | 4721701 | 645.5 | 2.657% | 2.981% | 601.7 | 699.3 | FAIL
AIM 5 | Run3 | Panel1 | 12 / 12 | 346676 | 44.6 | 3.465% | 3.408% | 41.4 | 48.2 | PASS
AIM 5 | Run3 | Panel2 | 12 / 12 | 1218773 | 158.3 | 2.666% | 3.333% | 145.6 | 169.2 | PASS
AIM 5 | Run3 | Panel3 | 12 / 12 | 3450821 | 460.1 | 3.986% | 2.955% | 424.8 | 493.6 | FAIL
AIM 5 | Run3 | Panel4 | 12 / 12 | 7565846 | 1144 | 3.458% | 2.357% | 1047.7 | 1217.5 | FAIL
AIM 5 | Run3 | RefCon1 | 12 / 12 | 294024 | 37.7 | 1.774% | 1.88% | 34.1 | 39.7 | PASS
AIM 5 | Run3 | RefCon2 | 12 / 12 | 1198908 | 151.4 | 2.743% | 2.616% | 138.3 | 160.7 | PASS
AIM 5 | Run3 | RefCon3 | 12 / 12 | 4781626 | 652.8 | 2.114% | 1.784% | 601.7 | 699.3 | PASS
Lot1 — Between-Run Validity (RefCon Grand Mean)
Sample | Instrument | Run1 Mean | Run2 Mean | Run3 Mean | Grand Mean | SD | %CV | Mean Spec Low | Mean Spec High | Result
RefCon1 | AIM 4 | 37.1 | 36.8 | 37.3 | 37 | 0.24 | 0.65% | 34.1 | 39.7 | PASS
RefCon1 | AIM 5 | 37.3 | 36.6 | 37.7 | 37.2 | 0.54 | 1.45% | 34.1 | 39.7 | PASS
RefCon2 | AIM 4 | 150.6 | 150 | 149.1 | 149.9 | 0.78 | 0.52% | 138.3 | 160.7 | PASS
RefCon2 | AIM 5 | 148.5 | 150.1 | 151.4 | 150 | 1.45 | 0.97% | 138.3 | 160.7 | PASS
RefCon3 | AIM 4 | 648.3 | 649.5 | 653.4 | 650.4 | 2.65 | 0.41% | 601.7 | 699.3 | PASS
RefCon3 | AIM 5 | 653.5 | 645.5 | 652.8 | 650.6 | 4.43 | 0.68% | 601.7 | 699.3 | PASS
Lot2 — Run Summary
Instrument | Run ID | Sample | Reps | RLU Avg | Dose Avg (pmol/L) | RLU %CV | Dose %CV | Mean Spec Low | Mean Spec High | Result
AIM 4 | Run1 | Panel1 | 12 / 12 | 344584 | 44.2 | 3.382% | 2.717% | 41.4 | 48.2 | PASS
AIM 4 | Run1 | Panel2 | 12 / 12 | 1206546 | 156.9 | 2.827% | 2.464% | 145.6 | 169.2 | PASS
AIM 4 | Run1 | Panel3 | 12 / 12 | 3475818 | 460.7 | 2.832% | 3.14% | 424.8 | 493.6 | PASS
AIM 4 | Run1 | Panel4 | 12 / 12 | 7456945 | 1128.2 | 3.366% | 2.756% | 1047.7 | 1217.5 | FAIL
AIM 4 | Run1 | RefCon1 | 12 / 12 | 288082 | 36.9 | 3.134% | 3.139% | 34.1 | 39.7 | FAIL
AIM 4 | Run1 | RefCon2 | 12 / 12 | 1179594 | 149.6 | 2.742% | 2.127% | 138.3 | 160.7 | PASS
AIM 4 | Run1 | RefCon3 | 12 / 12 | 4806046 | 658.2 | 2.925% | 2.519% | 601.7 | 699.3 | FAIL
AIM 4 | Run2 | Panel1 | 12 / 12 | 345180 | 44.2 | 5.07% | 5.378% | 41.4 | 48.2 | FAIL
AIM 4 | Run2 | Panel2 | 12 / 12 | 1204440 | 156.6 | 3.198% | 3.397% | 145.6 | 169.2 | FAIL
AIM 4 | Run2 | Panel3 | 12 / 12 | 3434408 | 460.1 | 3.336% | 3.228% | 424.8 | 493.6 | FAIL
AIM 4 | Run2 | Panel4 | 12 / 12 | 7452378 | 1133.4 | 2.379% | 1.477% | 1047.7 | 1217.5 | PASS
AIM 4 | Run2 | RefCon1 | 12 / 12 | 290220 | 37.2 | 2.376% | 2.025% | 34.1 | 39.7 | PASS
AIM 4 | Run2 | RefCon2 | 12 / 12 | 1199185 | 152.1 | 2.441% | 2.301% | 138.3 | 160.7 | PASS
AIM 4 | Run2 | RefCon3 | 12 / 12 | 4771935 | 650.4 | 2.441% | 2.057% | 601.7 | 699.3 | PASS
AIM 4 | Run3 | Panel1 | 12 / 12 | 347627 | 44.8 | 5.588% | 5.13% | 41.4 | 48.2 | FAIL
AIM 4 | Run3 | Panel2 | 12 / 12 | 1199762 | 156.5 | 3.38% | 3.585% | 145.6 | 169.2 | FAIL
AIM 4 | Run3 | Panel3 | 12 / 12 | 3418081 | 454.9 | 2.738% | 2.244% | 424.8 | 493.6 | PASS
AIM 4 | Run3 | Panel4 | 12 / 12 | 7435126 | 1129.2 | 2.808% | 2.077% | 1047.7 | 1217.5 | FAIL
AIM 4 | Run3 | RefCon1 | 12 / 12 | 291238 | 37.2 | 4.431% | 4.353% | 34.1 | 39.7 | FAIL
AIM 4 | Run3 | RefCon2 | 12 / 12 | 1172510 | 149.1 | 3.287% | 3.297% | 138.3 | 160.7 | FAIL
AIM 4 | Run3 | RefCon3 | 12 / 12 | 4738527 | 649.7 | 2.622% | 2.815% | 601.7 | 699.3 | FAIL
AIM 5 | Run1 | Panel1 | 12 / 12 | 345590 | 44.4 | 1.357% | 2.401% | 41.4 | 48.2 | PASS
AIM 5 | Run1 | Panel2 | 12 / 12 | 1219936 | 158.1 | 4.026% | 3.617% | 145.6 | 169.2 | FAIL
AIM 5 | Run1 | Panel3 | 12 / 12 | 3509140 | 465.9 | 2.933% | 2.884% | 424.8 | 493.6 | FAIL
AIM 5 | Run1 | Panel4 | 12 / 12 | 7414599 | 1124.8 | 2.61% | 2.694% | 1047.7 | 1217.5 | FAIL
AIM 5 | Run1 | RefCon1 | 12 / 12 | 285664 | 36.5 | 3.885% | 3.473% | 34.1 | 39.7 | FAIL
AIM 5 | Run1 | RefCon2 | 12 / 12 | 1201688 | 151.8 | 3.155% | 3.008% | 138.3 | 160.7 | FAIL
AIM 5 | Run1 | RefCon3 | 12 / 12 | 4703724 | 645.3 | 3.06% | 3.115% | 601.7 | 699.3 | FAIL
AIM 5 | Run2 | Panel1 | 12 / 12 | 351366 | 45.1 | 3.832% | 3.73% | 41.4 | 48.2 | FAIL
AIM 5 | Run2 | Panel2 | 12 / 12 | 1209972 | 157.6 | 2.99% | 2.853% | 145.6 | 169.2 | PASS
AIM 5 | Run2 | Panel3 | 12 / 12 | 3452241 | 460.9 | 2.475% | 2.382% | 424.8 | 493.6 | PASS
AIM 5 | Run2 | Panel4 | 12 / 12 | 7408633 | 1129.4 | 1.64% | 1.745% | 1047.7 | 1217.5 | PASS
AIM 5 | Run2 | RefCon1 | 12 / 12 | 289301 | 37 | 2.445% | 2.609% | 34.1 | 39.7 | PASS
AIM 5 | Run2 | RefCon2 | 12 / 12 | 1207739 | 152.7 | 3.598% | 2.86% | 138.3 | 160.7 | FAIL
AIM 5 | Run2 | RefCon3 | 12 / 12 | 4736304 | 645.6 | 2.155% | 2.16% | 601.7 | 699.3 | PASS
AIM 5 | Run3 | Panel1 | 12 / 12 | 346057 | 44.3 | 4.315% | 4.716% | 41.4 | 48.2 | FAIL
AIM 5 | Run3 | Panel2 | 12 / 12 | 1207738 | 156.2 | 3.034% | 3.223% | 145.6 | 169.2 | PASS
AIM 5 | Run3 | Panel3 | 12 / 12 | 3472707 | 462.7 | 2.27% | 1.954% | 424.8 | 493.6 | PASS
AIM 5 | Run3 | Panel4 | 12 / 12 | 7500385 | 1137.5 | 2.572% | 2.285% | 1047.7 | 1217.5 | FAIL
AIM 5 | Run3 | RefCon1 | 12 / 12 | 283394 | 36.5 | 3.572% | 3.333% | 34.1 | 39.7 | FAIL
AIM 5 | Run3 | RefCon2 | 12 / 12 | 1177554 | 148.9 | 3.719% | 2.807% | 138.3 | 160.7 | FAIL
AIM 5 | Run3 | RefCon3 | 12 / 12 | 4790135 | 659.2 | 2.029% | 1.746% | 601.7 | 699.3 | PASS
Lot2 — Between-Run Validity (RefCon Grand Mean)
Sample | Instrument | Run1 Mean | Run2 Mean | Run3 Mean | Grand Mean | SD | %CV | Mean Spec Low | Mean Spec High | Result
RefCon1 | AIM 4 | 36.9 | 37.2 | 37.2 | 37.1 | 0.2 | 0.54% | 34.1 | 39.7 | PASS
RefCon1 | AIM 5 | 36.5 | 37 | 36.5 | 36.7 | 0.29 | 0.79% | 34.1 | 39.7 | PASS
RefCon2 | AIM 4 | 149.6 | 152.1 | 149.1 | 150.3 | 1.57 | 1.04% | 138.3 | 160.7 | PASS
RefCon2 | AIM 5 | 151.8 | 152.7 | 148.9 | 151.1 | 2 | 1.32% | 138.3 | 160.7 | PASS
RefCon3 | AIM 4 | 658.2 | 650.4 | 649.7 | 652.8 | 4.71 | 0.72% | 601.7 | 699.3 | PASS
RefCon3 | AIM 5 | 645.3 | 645.6 | 659.2 | 650.1 | 7.91 | 1.22% | 601.7 | 699.3 | PASS
Validity — Single-Replicate RefCon Runs
Instrument | Run ID | Sample | RLU | Result (pmol/L) | Spec Low | Spec High | Result
AIM 4 | Run1 | RefCon1 | 287515 | 36.4 | 28.6 | 45.2 | PASS
AIM 4 | Run1 | RefCon2 | 1172630 | 149.6 | 115.9 | 183.1 | PASS
AIM 4 | Run1 | RefCon3 | 4776332 | 661.5 | 504.1 | 796.9 | PASS
AIM 4 | Run2 | RefCon1 | 289802 | 37.6 | 28.6 | 45.2 | PASS
AIM 4 | Run2 | RefCon2 | 1208485 | 150.4 | 115.9 | 183.1 | PASS
AIM 4 | Run2 | RefCon3 | 4764968 | 647.6 | 504.1 | 796.9 | PASS
AIM 4 | Run3 | RefCon1 | 291011 | 37.2 | 28.6 | 45.2 | PASS
AIM 4 | Run3 | RefCon2 | 1196059 | 151.9 | 115.9 | 183.1 | PASS
AIM 4 | Run3 | RefCon3 | 4546657 | 635 | 504.1 | 796.9 | PASS
AIM 5 | Run1 | RefCon1 | 288635 | 36.8 | 28.6 | 45.2 | PASS
AIM 5 | Run1 | RefCon2 | 1191027 | 150 | 115.9 | 183.1 | PASS
AIM 5 | Run1 | RefCon3 | 4770718 | 641.8 | 504.1 | 796.9 | PASS
AIM 5 | Run2 | RefCon1 | 294496 | 37.2 | 28.6 | 45.2 | PASS
AIM 5 | Run2 | RefCon2 | 1175052 | 147.2 | 115.9 | 183.1 | PASS
AIM 5 | Run2 | RefCon3 | 4819829 | 658.7 | 504.1 | 796.9 | PASS
AIM 5 | Run3 | RefCon1 | 300456 | 38.1 | 28.6 | 45.2 | PASS
AIM 5 | Run3 | RefCon2 | 1180531 | 147.6 | 115.9 | 183.1 | PASS
AIM 5 | Run3 | RefCon3 | 4741321 | 653.4 | 504.1 | 796.9 | PASS
Validity — Grand Mean Acceptance
Sample | # Values | Grand Mean (pmol/L) | Spec Low | Spec High | Min | Max | Result
RefCon1 | 6 | 37.2 | 28.6 | 45.2 | 36.4 | 38.1 | PASS
RefCon2 | 6 | 149.4 | 115.9 | 183.1 | 147.2 | 151.9 | PASS
RefCon3 | 6 | 649.7 | 504.1 | 796.9 | 635 | 661.5 | PASS

## Panel1
Panel1 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 349386 | 45 | 2.75% | 34.7 | 54.9 | 41.4 | 48.2 | PASS
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 349485 | 44.9 | 3.836% | 34.7 | 54.9 | 41.4 | 48.2 | FAIL
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 343130 | 43.8 | 2.397% | 34.7 | 54.9 | 41.4 | 48.2 | PASS
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 343362 | 43.9 | 1.898% | 34.7 | 54.9 | 41.4 | 48.2 | PASS
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 347408 | 44.4 | 3.866% | 34.7 | 54.9 | 41.4 | 48.2 | FAIL
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 346676 | 44.6 | 3.408% | 34.7 | 54.9 | 41.4 | 48.2 | PASS
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 344584 | 44.2 | 2.717% | 34.7 | 54.9 | 41.4 | 48.2 | PASS
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 345180 | 44.2 | 5.378% | 34.7 | 54.9 | 41.4 | 48.2 | FAIL
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 347627 | 44.8 | 5.13% | 34.7 | 54.9 | 41.4 | 48.2 | FAIL
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 345590 | 44.4 | 2.401% | 34.7 | 54.9 | 41.4 | 48.2 | PASS
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 351366 | 45.1 | 3.73% | 34.7 | 54.9 | 41.4 | 48.2 | FAIL
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 346057 | 44.3 | 4.716% | 34.7 | 54.9 | 41.4 | 48.2 | FAIL
Lot1 — Individual Replicate Values (Indiv Spec: 34.7–54.9 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 45.3 | 44.1 | 43.1 | 43.4 | 43.9 | 44.4
2 | 44.3 | 43.3 | 44 | 43.7 | 43.9 | 45.6
3 | 45.6 | 43 | 45.2 | 43.5 | 42.2 | 42.7
4 | 47 | 43.2 | 43.6 | 43.1 | 43.2 | 45.1
5 | 44.1 | 44.4 | 43.1 | 44.6 | 42.7 | 44.8
6 | 44.1 | 45.1 | 44.9 | 43.1 | 47.4 | 45.3
7 | 47.1 | 45 | 44.9 | 44.9 | 46 | 45
8 | 45.8 | 45.9 | 43.7 | 44.4 | 46.6 | 48.5
9 | 43.7 | 44.5 | 43.7 | 44.1 | 45.7 | 43.4
10 | 45.4 | 46.9 | 44.9 | 43 | 42.6 | 43.6
11 | 43.7 | 44.1 | 42.1 | 43.6 | 43.6 | 43.5
12 | 43.7 | 49 | 42.2 | 45.7 | 45.3 | 43.6
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 1.41
Mean | AIM 4 Run1: 44.4
%CV | AIM 4 Run1: 3.2%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 34.7–54.9 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 43.2 | 48 | 47.9 | 45.9 | 45.3 | 46.5
2 | 43.7 | 46.2 | 47.4 | 43.1 | 44.8 | 42
3 | 46 | 42 | 44.1 | 44.6 | 43.1 | 43.6
4 | 45 | 43.7 | 46.1 | 43.7 | 45.7 | 45.1
5 | 43.6 | 46.6 | 45.6 | 45.3 | 43.6 | 47.1
6 | 45.3 | 43.3 | 46.8 | 45 | 44.7 | 44.4
7 | 44.7 | 45.2 | 42.9 | 46.2 | 48.7 | 43.6
8 | 46.1 | 45.8 | 45.6 | 43.7 | 44.3 | 47.6
9 | 43.3 | 43 | 46.2 | 44.1 | 46.4 | 42.1
10 | 44 | 44.4 | 41.6 | 42.9 | 43.3 | 40.9
11 | 43.9 | 39.2 | 42.6 | 43.8 | 44.4 | 45.2
12 | 42.1 | 42.8 | 41.1 | 45.1 | 47.4 | 43.7
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 1.83
Mean | AIM 4 Run1: 44.5
%CV | AIM 4 Run1: 4.1%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES

## Panel2
Panel2 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 1183139 | 154.1 | 3.149% | 122 | 192.8 | 145.6 | 169.2 | FAIL
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 1213263 | 156.2 | 2.477% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 1205130 | 156.9 | 2.191% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 1197396 | 155 | 2.999% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 1219369 | 157.3 | 2.351% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 1218773 | 158.3 | 3.333% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 1206546 | 156.9 | 2.464% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 1204440 | 156.6 | 3.397% | 122 | 192.8 | 145.6 | 169.2 | FAIL
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 1199762 | 156.5 | 3.585% | 122 | 192.8 | 145.6 | 169.2 | FAIL
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 1219936 | 158.1 | 3.617% | 122 | 192.8 | 145.6 | 169.2 | FAIL
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 1209972 | 157.6 | 2.853% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 1207738 | 156.2 | 3.223% | 122 | 192.8 | 145.6 | 169.2 | PASS
Lot1 — Individual Replicate Values (Indiv Spec: 122.0–192.8 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 158.2 | 160 | 153.5 | 159.4 | 151.1 | 153.9
2 | 147.7 | 152.8 | 156 | 152.2 | 160.5 | 162.8
3 | 148.6 | 151.8 | 158.5 | 157.5 | 155.8 | 163.9
4 | 154.3 | 159.3 | 164.2 | 160.7 | 155.2 | 154.2
5 | 152.1 | 155.9 | 161.2 | 148.9 | 160.5 | 152.9
6 | 158.5 | 160.5 | 156.2 | 159.6 | 159.2 | 159.3
7 | 152.6 | 159.3 | 156.9 | 153.8 | 155.2 | 154.3
8 | 150.1 | 156.6 | 152.1 | 159.8 | 162.6 | 160.1
9 | 164.1 | 152.9 | 156.9 | 153.3 | 151.7 | 158
10 | 155.9 | 149.6 | 155.6 | 148.2 | 160 | 149.6
11 | 157.3 | 154.8 | 158.6 | 149.1 | 159.9 | 164.5
12 | 150.1 | 161.2 | 153 | 157.2 | 155.5 | 165.7
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 4.42
Mean | AIM 4 Run1: 156.3
%CV | AIM 4 Run1: 2.8%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 122.0–192.8 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 158.4 | 155.8 | 155.7 | 160.7 | 153.9 | 152
2 | 158.3 | 150.9 | 160.5 | 152.5 | 165.8 | 160.4
3 | 157 | 164.9 | 164.3 | 161.2 | 160.4 | 158.2
4 | 155.9 | 150 | 157.4 | 163.6 | 154.3 | 154.3
5 | 150.1 | 154.9 | 164.9 | 159 | 160.1 | 150.8
6 | 155 | 157.6 | 150.3 | 166.1 | 161.7 | 161.2
7 | 155.3 | 164 | 148.7 | 153.2 | 160 | 160.2
8 | 153.1 | 150 | 156.7 | 150.9 | 149.4 | 156.5
9 | 156.2 | 162.7 | 158.9 | 148.3 | 153.5 | 166
10 | 159 | 157 | 156.8 | 164.3 | 155.8 | 151.8
11 | 166.2 | 152.2 | 146.9 | 160.2 | 156.6 | 149.6
12 | 157.8 | 159.2 | 156.6 | 156.7 | 160 | 153.6
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 4.91
Mean | AIM 4 Run1: 157
%CV | AIM 4 Run1: 3.1%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES

## Panel3
Panel3 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 3433440 | 456.4 | 2.618% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 3446366 | 459.8 | 2.237% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 3456700 | 462 | 1.674% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 3411116 | 454.2 | 2.27% | 355.9 | 562.5 | 424.8 | 493.6 | FAIL
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 3413381 | 456.5 | 3.232% | 355.9 | 562.5 | 424.8 | 493.6 | FAIL
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 3450821 | 460.1 | 2.955% | 355.9 | 562.5 | 424.8 | 493.6 | FAIL
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 3475818 | 460.7 | 3.14% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 3434408 | 460.1 | 3.228% | 355.9 | 562.5 | 424.8 | 493.6 | FAIL
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 3418081 | 454.9 | 2.244% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 3509140 | 465.9 | 2.884% | 355.9 | 562.5 | 424.8 | 493.6 | FAIL
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 3452241 | 460.9 | 2.382% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 3472707 | 462.7 | 1.954% | 355.9 | 562.5 | 424.8 | 493.6 | PASS
Lot1 — Individual Replicate Values (Indiv Spec: 355.9–562.5 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 451.8 | 461.8 | 465.9 | 462.5 | 463.3 | 450.8
2 | 460.5 | 442.4 | 479.4 | 447 | 442.3 | 453.8
3 | 443.7 | 461.3 | 457.6 | 467.5 | 471.3 | 462.8
4 | 464 | 464.1 | 464.3 | 436.9 | 456.5 | 463.5
5 | 451 | 447.2 | 468.2 | 458.1 | 452 | 467.8
6 | 455.1 | 461 | 453.7 | 442.9 | 473 | 485.8
7 | 451 | 459.8 | 462 | 450.3 | 449.6 | 456.6
8 | 483.7 | 443.8 | 459.2 | 459.6 | 440.3 | 448.4
9 | 458.8 | 463.8 | 460.3 | 447.5 | 438.3 | 440.6
10 | 444.9 | 466.5 | 448.7 | 453.9 | 467.1 | 449.3
11 | 469.9 | 473.4 | 459.3 | 472.4 | 442 | 458.6
12 | 442.7 | 473 | 465.6 | 451.3 | 482.4 | 482.9
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 11.57
Mean | AIM 4 Run1: 458.2
%CV | AIM 4 Run1: 2.5%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 355.9–562.5 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 462.4 | 461.6 | 441.6 | 462.7 | 461.4 | 458.4
2 | 458 | 451 | 467.9 | 444 | 441.2 | 462.2
3 | 433.5 | 459.9 | 463.9 | 491.6 | 464.1 | 455.8
4 | 458.6 | 453.9 | 446.5 | 460.7 | 467.1 | 463.7
5 | 459.8 | 460.5 | 452.2 | 460.5 | 466.5 | 442.3
6 | 491.8 | 467.8 | 444.9 | 468.7 | 473.4 | 478.2
7 | 456.4 | 480.1 | 458.2 | 465.4 | 470.1 | 457.9
8 | 463 | 442.5 | 471.7 | 462 | 465.1 | 473.9
9 | 458.5 | 487.4 | 445.9 | 448.5 | 458.1 | 463.6
10 | 443.4 | 433 | 465.7 | 465.3 | 436.9 | 465.1
11 | 474.2 | 457 | 451.9 | 484.1 | 464.7 | 466.6
12 | 469 | 466.8 | 448.4 | 476.9 | 461.8 | 465
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 12.37
Mean | AIM 4 Run1: 460.9
%CV | AIM 4 Run1: 2.7%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES

## Panel4
Panel4 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 7369988 | 1121 | 2.196% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 7523870 | 1145.5 | 3.042% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 7478332 | 1134.9 | 3.22% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 7415080 | 1126.6 | 2.493% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 7464346 | 1131.1 | 2.18% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 7565846 | 1144 | 2.357% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 7456945 | 1128.2 | 2.756% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 7452378 | 1133.4 | 1.477% | 877.8 | 1387.4 | 1047.7 | 1217.5 | PASS
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 7435126 | 1129.2 | 2.077% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 7414599 | 1124.8 | 2.694% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 7408633 | 1129.4 | 1.745% | 877.8 | 1387.4 | 1047.7 | 1217.5 | PASS
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 7500385 | 1137.5 | 2.285% | 877.8 | 1387.4 | 1047.7 | 1217.5 | FAIL
Lot1 — Individual Replicate Values (Indiv Spec: 877.8–1387.4 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 1135.7 | 1092.6 | 1169.4 | 1152.7 | 1073.5 | 1116
2 | 1076.9 | 1104.6 | 1156 | 1099.4 | 1176 | 1136.1
3 | 1094 | 1144 | 1188.4 | 1144.4 | 1135.7 | 1129.6
4 | 1135.3 | 1143.9 | 1109.2 | 1169.1 | 1127.4 | 1162.2
5 | 1150 | 1144 | 1153.7 | 1063 | 1115.2 | 1198.5
6 | 1134.6 | 1234.5 | 1135 | 1108.4 | 1140.8 | 1115.6
7 | 1126.9 | 1145.5 | 1189.4 | 1145.7 | 1129 | 1116.7
8 | 1121.8 | 1160.8 | 1108.1 | 1124.5 | 1159.9 | 1158.3
9 | 1089.9 | 1155.9 | 1107.2 | 1140.1 | 1133.1 | 1148.5
10 | 1110.5 | 1147.7 | 1113.7 | 1113.6 | 1134.1 | 1180.1
11 | 1117.5 | 1121.4 | 1072.4 | 1132.3 | 1120.1 | 1145.8
12 | 1158.7 | 1150.6 | 1115.7 | 1125.8 | 1128.5 | 1120.3
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 29.94
Mean | AIM 4 Run1: 1133.8
%CV | AIM 4 Run1: 2.6%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 877.8–1387.4 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 1151.5 | 1137.6 | 1127.1 | 1173.2 | 1137.4 | 1147.4
2 | 1105.3 | 1113.1 | 1101.9 | 1116.1 | 1095.4 | 1166
3 | 1168 | 1124.4 | 1115 | 1103.2 | 1100.7 | 1135.3
4 | 1092 | 1116.6 | 1097.5 | 1126.6 | 1158.6 | 1149.2
5 | 1145.9 | 1114 | 1183.3 | 1131.5 | 1128.9 | 1127.6
6 | 1189.4 | 1153 | 1131 | 1159.7 | 1148.5 | 1169.1
7 | 1103.1 | 1139.7 | 1111 | 1084.1 | 1130.8 | 1111.7
8 | 1114.6 | 1111.2 | 1135.8 | 1171.5 | 1130.8 | 1178.8
9 | 1132.7 | 1154.4 | 1127 | 1125.7 | 1155.4 | 1128.9
10 | 1116.3 | 1138.3 | 1124 | 1118.4 | 1116 | 1091.2
11 | 1087.9 | 1152 | 1146.7 | 1102.6 | 1132.6 | 1133.5
12 | 1131.9 | 1147.1 | 1150.5 | 1085.1 | 1117.5 | 1111.5
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 24.54
Mean | AIM 4 Run1: 1130.4
%CV | AIM 4 Run1: 2.2%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES

## RefCon1
RefCon1 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 288587 | 37.1 | 2.579% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 288467 | 36.8 | 3.556% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 291894 | 37.3 | 3.176% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 290542 | 37.3 | 2.365% | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 286083 | 36.6 | 2.285% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 294024 | 37.7 | 1.88% | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 288082 | 36.9 | 3.139% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 290220 | 37.2 | 2.025% | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 291238 | 37.2 | 4.353% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 285664 | 36.5 | 3.473% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 289301 | 37 | 2.609% | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 283394 | 36.5 | 3.333% | 28.6 | 45.2 | 34.1 | 39.7 | FAIL
Validity | Validity | AIM 4 | Run1 | 1 / 1 | 287515 | 36.4 |  | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Validity | Validity | AIM 4 | Run2 | 1 / 1 | 289802 | 37.6 |  | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Validity | Validity | AIM 4 | Run3 | 1 / 1 | 291011 | 37.2 |  | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Validity | Validity | AIM 5 | Run1 | 1 / 1 | 288635 | 36.8 |  | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Validity | Validity | AIM 5 | Run2 | 1 / 1 | 294496 | 37.2 |  | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Validity | Validity | AIM 5 | Run3 | 1 / 1 | 300456 | 38.1 |  | 28.6 | 45.2 | 34.1 | 39.7 | PASS
Lot1 — Individual Replicate Values (Indiv Spec: 28.6–45.2 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 37.4 | 36.1 | 36.2 | 38.3 | 37.3 | 37.7
2 | 35 | 36.7 | 37.2 | 37.3 | 35.1 | 38.2
3 | 37.4 | 36.5 | 37.4 | 37.4 | 35.5 | 37.9
4 | 36.6 | 37.1 | 39.1 | 36.5 | 37.8 | 37.6
5 | 36.2 | 39.6 | 38.1 | 36.5 | 37.2 | 38.2
6 | 37.7 | 34.9 | 36.4 | 36.5 | 36.8 | 38.3
7 | 38.1 | 37.8 | 36 | 37.4 | 37 | 38.5
8 | 38 | 35.2 | 37.5 | 36.5 | 37.4 | 37.7
9 | 36.1 | 36.5 | 35.5 | 37.3 | 36.4 | 36.8
10 | 36.7 | 38.2 | 39 | 39.3 | 36.1 | 37.2
11 | 37.4 | 37.1 | 38.3 | 38 | 37.2 | 38.3
12 | 38.1 | 35.8 | 36.5 | 36.6 | 35.9 | 36.1
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 1.03
Mean | AIM 4 Run1: 37.1
%CV | AIM 4 Run1: 2.8%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 28.6–45.2 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 35.8 | 36.1 | 36.4 | 37.9 | 36.5 | 37.9
2 | 37.5 | 36.4 | 36.4 | 37.1 | 36.7 | 36.3
3 | 36 | 37.8 | 36.7 | 35.6 | 37.2 | 36.5
4 | 38.7 | 37.7 | 34.4 | 35.6 | 36.5 | 34.9
5 | 36.1 | 37 | 35.3 | 36.6 | 38.4 | 36.5
6 | 36.6 | 37.1 | 38.5 | 38.9 | 36 | 34.3
7 | 37.9 | 38.4 | 38.8 | 36.7 | 36.8 | 35.2
8 | 35.6 | 36.3 | 36.7 | 35.3 | 36.5 | 37.8
9 | 37.3 | 37.6 | 37.6 | 36.7 | 38.6 | 37.9
10 | 38.5 | 36.8 | 37.3 | 36.7 | 37.2 | 37.5
11 | 35.2 | 36.8 | 40.4 | 34 | 38.1 | 35.9
12 | 37.2 | 38.2 | 38.2 | 36.9 | 35.4 | 36.9
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 1.19
Mean | AIM 4 Run1: 36.9
%CV | AIM 4 Run1: 3.2%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Validity — Single-Replicate Values (Spec: 28.6–45.2 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 36.4 | 37.6 | 37.2 | 36.8 | 37.2 | 38.1
Accepted | YES | YES | YES | YES | YES | YES

## RefCon2
RefCon2 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 1192100 | 150.6 | 2.736% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 1180225 | 150 | 3.163% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 1184751 | 149.1 | 2.95% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 1177697 | 148.5 | 3.83% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 1185646 | 150.1 | 3.458% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 1198908 | 151.4 | 2.616% | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 1179594 | 149.6 | 2.127% | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 1199185 | 152.1 | 2.301% | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 1172510 | 149.1 | 3.297% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 1201688 | 151.8 | 3.008% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 1207739 | 152.7 | 2.86% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 1177554 | 148.9 | 2.807% | 115.9 | 183.1 | 138.3 | 160.7 | FAIL
Validity | Validity | AIM 4 | Run1 | 1 / 1 | 1172630 | 149.6 |  | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Validity | Validity | AIM 4 | Run2 | 1 / 1 | 1208485 | 150.4 |  | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Validity | Validity | AIM 4 | Run3 | 1 / 1 | 1196059 | 151.9 |  | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Validity | Validity | AIM 5 | Run1 | 1 / 1 | 1191027 | 150 |  | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Validity | Validity | AIM 5 | Run2 | 1 / 1 | 1175052 | 147.2 |  | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Validity | Validity | AIM 5 | Run3 | 1 / 1 | 1180531 | 147.6 |  | 115.9 | 183.1 | 138.3 | 160.7 | PASS
Lot1 — Individual Replicate Values (Indiv Spec: 115.9–183.1 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 147.8 | 146.8 | 142.3 | 155.4 | 151.8 | 151.7
2 | 149.2 | 153.1 | 156.1 | 148.2 | 142.3 | 148.2
3 | 145 | 146.7 | 149.5 | 140.8 | 154.6 | 150.1
4 | 144.6 | 151 | 155.6 | 145.5 | 152.1 | 155.8
5 | 153.7 | 150.2 | 142.8 | 141.6 | 151.2 | 150.9
6 | 156.1 | 147.1 | 147.3 | 148.4 | 154.4 | 150.2
7 | 149.7 | 159.6 | 150 | 150.1 | 157.5 | 143.9
8 | 154.5 | 152.9 | 150.2 | 157.5 | 154.6 | 153.4
9 | 151.6 | 140.9 | 148 | 151.5 | 141.7 | 152.9
10 | 147.1 | 150.8 | 152.8 | 149 | 144.2 | 159.7
11 | 151.6 | 147 | 145.2 | 153.7 | 147.2 | 148.6
12 | 156.9 | 153.8 | 149.4 | 140 | 150.1 | 151
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 4.65
Mean | AIM 4 Run1: 150
%CV | AIM 4 Run1: 3.1%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 115.9–183.1 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 151.2 | 153.7 | 149.4 | 149 | 151.2 | 150
2 | 153.5 | 153.7 | 145.7 | 153.1 | 154 | 144.8
3 | 144.4 | 155.9 | 142.8 | 158.3 | 150.4 | 156.8
4 | 144.1 | 150.1 | 150.9 | 155.1 | 154.8 | 153.9
5 | 152.3 | 153.1 | 146.6 | 148.8 | 147.7 | 149
6 | 151.3 | 148.6 | 143.6 | 145 | 156.3 | 150.1
7 | 151.1 | 151.5 | 147.1 | 161.6 | 160.3 | 150.9
8 | 151.6 | 149.4 | 145.1 | 150.3 | 148.4 | 140.8
9 | 146.9 | 150.4 | 157.6 | 150.1 | 148 | 148.9
10 | 151 | 152.7 | 154 | 149.9 | 156.5 | 146.9
11 | 151.3 | 146.3 | 150 | 150.9 | 157.1 | 145.5
12 | 146.8 | 159.4 | 156.7 | 149.4 | 147.6 | 148.7
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 4.3
Mean | AIM 4 Run1: 150.7
%CV | AIM 4 Run1: 2.9%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Validity — Single-Replicate Values (Spec: 115.9–183.1 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 149.6 | 150.4 | 151.9 | 150 | 147.2 | 147.6
Accepted | YES | YES | YES | YES | YES | YES

## RefCon3
RefCon3 — TP1 Run Summary (All Lots)
Lot | Lot Type | Instrument | Run ID | Reps | RLU Avg | Dose Avg (pmol/L) | Dose %CV | Indiv Rep Spec Low | Indiv Rep Spec High | Mean Spec Low | Mean Spec High | Pass/Fail
Lot1 | Feasibility | AIM 4 | Run1 | 12 / 12 | 4738966 | 648.3 | 3.074% | 504.1 | 796.9 | 601.7 | 699.3 | FAIL
Lot1 | Feasibility | AIM 4 | Run2 | 12 / 12 | 4708026 | 649.5 | 2.332% | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Lot1 | Feasibility | AIM 4 | Run3 | 12 / 12 | 4763350 | 653.4 | 2.375% | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Lot1 | Feasibility | AIM 5 | Run1 | 12 / 12 | 4785178 | 653.5 | 2.422% | 504.1 | 796.9 | 601.7 | 699.3 | FAIL
Lot1 | Feasibility | AIM 5 | Run2 | 12 / 12 | 4721701 | 645.5 | 2.981% | 504.1 | 796.9 | 601.7 | 699.3 | FAIL
Lot1 | Feasibility | AIM 5 | Run3 | 12 / 12 | 4781626 | 652.8 | 1.784% | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Lot2 | Feasibility | AIM 4 | Run1 | 12 / 12 | 4806046 | 658.2 | 2.519% | 504.1 | 796.9 | 601.7 | 699.3 | FAIL
Lot2 | Feasibility | AIM 4 | Run2 | 12 / 12 | 4771935 | 650.4 | 2.057% | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Lot2 | Feasibility | AIM 4 | Run3 | 12 / 12 | 4738527 | 649.7 | 2.815% | 504.1 | 796.9 | 601.7 | 699.3 | FAIL
Lot2 | Feasibility | AIM 5 | Run1 | 12 / 12 | 4703724 | 645.3 | 3.115% | 504.1 | 796.9 | 601.7 | 699.3 | FAIL
Lot2 | Feasibility | AIM 5 | Run2 | 12 / 12 | 4736304 | 645.6 | 2.16% | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Lot2 | Feasibility | AIM 5 | Run3 | 12 / 12 | 4790135 | 659.2 | 1.746% | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Validity | Validity | AIM 4 | Run1 | 1 / 1 | 4776332 | 661.5 |  | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Validity | Validity | AIM 4 | Run2 | 1 / 1 | 4764968 | 647.6 |  | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Validity | Validity | AIM 4 | Run3 | 1 / 1 | 4546657 | 635 |  | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Validity | Validity | AIM 5 | Run1 | 1 / 1 | 4770718 | 641.8 |  | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Validity | Validity | AIM 5 | Run2 | 1 / 1 | 4819829 | 658.7 |  | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Validity | Validity | AIM 5 | Run3 | 1 / 1 | 4741321 | 653.4 |  | 504.1 | 796.9 | 601.7 | 699.3 | PASS
Lot1 — Individual Replicate Values (Indiv Spec: 504.1–796.9 pmol/L)
Rep | AIM 4 |  |  | AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 649.4 | 637.1 | 652 | 653.8 | 658.4 | 654.1
2 | 675.4 | 648.1 | 658.4 | 662.5 | 638.2 | 675.6
3 | 607.4 | 658.2 | 661.6 | 626 | 653 | 648.5
4 | 663.4 | 664.1 | 631.7 | 668.6 | 637.7 | 654.5
5 | 651.4 | 630.5 | 625.1 | 655.5 | 640.1 | 659.9
6 | 645.1 | 644.6 | 670.8 | 643.3 | 627.1 | 653
7 | 651.5 | 642.3 | 655.4 | 660.3 | 635 | 642.7
8 | 617.7 | 639.4 | 637.8 | 686.9 | 628 | 653.2
9 | 646.4 | 678.7 | 675.2 | 653 | 634.1 | 667.4
10 | 655.8 | 656.6 | 651.9 | 654 | 667.1 | 633.3
11 | 674 | 629.5 | 669.2 | 642.5 | 634.6 | 652.2
12 | 641.6 | 664.9 | 651.1 | 636.2 | 692.8 | 638.6
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 16.14
Mean | AIM 4 Run1: 650.5
%CV | AIM 4 Run1: 2.5%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Lot2 — Individual Replicate Values (Indiv Spec: 504.1–796.9 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 680.3 | 633.7 | 651.3 | 640.7 | 643.2 | 679.2
2 | 657.7 | 630.3 | 636 | 641.1 | 645.4 | 660.4
3 | 630.6 | 668.8 | 674.8 | 649.5 | 628.2 | 640.7
4 | 660.7 | 662.9 | 658.8 | 641.2 | 635.1 | 659.3
5 | 634.2 | 660.1 | 633.1 | 638.4 | 633.7 | 672.7
6 | 662.8 | 660.2 | 646.9 | 651.7 | 637.5 | 665
7 | 668.8 | 649.8 | 635.8 | 645.9 | 649.4 | 651
8 | 636.7 | 635.4 | 627.5 | 674.4 | 653.8 | 639.5
9 | 665.7 | 651.2 | 665.1 | 606.9 | 675.2 | 661.3
10 | 656.7 | 639 | 681 | 667.7 | 633.8 | 656.4
11 | 663.4 | 665.8 | 627.3 | 670.2 | 666 | 664.5
12 | 680.8 | 647.6 | 659.1 | 616.3 | 646.5 | 660.3
Accepted | YES | YES | YES | YES | YES | YES
Overall | bar: AIM 4 Run1
Std Dev | AIM 4 Run1: 16.31
Mean | AIM 4 Run1: 651.4
%CV | AIM 4 Run1: 2.5%
Valid? | AIM 4 Run1: Valid
Accepted | AIM 4 Run1: YES
Validity — Single-Replicate Values (Spec: 504.1–796.9 pmol/L)
Rep | AIM 4 Run1: AIM 4; AIM 5 Run1: AIM 5
| Run1 | Run2 | Run3 | Run1 | Run2 | Run3
1 | 661.5 | 647.6 | 635 | 641.8 | 658.7 | 653.4
Accepted | YES | YES | YES | YES | YES | YES